##### Taken from: http://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/ucm341718.htm#s4
##### 4.1.1 In vivo testing for SpO2 accuracy under laboratory conditions
We recommend you follow Clause 201.12.1.101.2 and Annex EE.2 of ISO 80601-2-61:2011 Procedure for invasive laboratory testing on healthy volunteers, or equivalent method to validate the SpO2 accuracy specifications of your pulse oximeter system by comparing each value from your system and a simultaneous value from co-oximetry of an arterial blood sample. We recommend you submit a detailed clinical report for this testing.  Your report should describe the device configuration tested and include the following:
- test apparatus used, including means for arterial catheterization and blood sampling, means for recording SpO2 values, and means for delivering medical grade oxygen-nitrogen mixtures of varying fractional inspired oxygen (FiO2) levels;
- inclusion and exclusion criteria;
- number of subjects;
- number of samples taken per subject;
- specific conditions of testing, including laboratory conditions, subject motion, low pulse amplitude;
- type and frequency of motion for testing under motion conditions, if applicable;
- criteria and methods for determining stability of reference arterial blood oxygen saturation (SaO2) at the pulse oximeter sensor site;
- desaturation profile, including target saturation plateaus and ranges; and
- formula used for determination of root mean square difference (Arms).  (See Clause 201.12.1.101.2.2 of ISO 80601-2-61:2011 for recommended formula.)
We recommend you conduct the study described in Clause 201.12.1.101.2 and Annex EE.2 of ISO 80601-2-61:2011 on 10 or more healthy subjects that vary in age and gender. Your data should include 200 or more data points (paired observations: pulse oximeter, co-oximeter). These data should be distributed as described in Annex EE.2.3.4(g).
Your study should have subjects with a range of skin pigmentations, including at least 2 darkly pigmented subjects or 15% of your subject pool, whichever is larger.
We recommend you provide a line listing, a Bland-Altman plot, error plots (i.e., SaO2 versus (SpO2-SaO2) for both individual test subjects and all subjects pooled), and rationale for any points excluded from analysis.  Please include population mean bias (μ0), between-subject variance (σμ2), within-subject variance (σ2), and upper 95% and lower 95% limits of agreement. Please provide this information as outlined in Section 3 of “Agreement Between Methods Of Measurement With Multiple Observations Per Individual” by Bland and Altman.2  If you note that the plots show noticeable outliers, please provide the following:
- a discussion of the state of health, subject characteristics, test setup, test procedure, and any other factors that may have affected these data points; and
- a discussion of how the outlier(s) do not raise safety and performance concerns regarding the accuracy of the device.
We recommend an Arms specification in conformance with Clause 201.12.1.101.1 of ISO 80601-2-61:2011. We recognize that accuracy is, among other things, a function of patient characteristics, application site and sensor geometry. The table below outlines the typical Arms between measured values (SpO2) and reference values (SaO2) under normal conditions ranging from 70% to 100% SpO2.