Calibration and validation of a high-quality open access pulse oximeter

Alexander Pavlosky1, Tarek Loubani1,2,3,4*

1 Faculty of Medicine, University of Western Ontario, London, Ontario, Canada.
2 Division of Emergency Medicine, Department of Medicine, University of Western Ontario, London, Ontario, Canada.
3 Division of Emergency Medicine, London Health Sciences Centre, London, Ontario, Canada
4 Division of Emergency Medicine, Al-Shifa Hospital, Gaza

* Corresponding author
E-mail: Tarek.Loubani@lhsc.on.ca

# Abstract

# Introduction

Pulse oximetry is an important tool used in high-income countries to assess a patient’s oxyhemoglobin saturation and is used diversely in medicine both in clinical and surgical settings. In the developed world, pulse oximetry has become the standard of care as it increases patient safety by alerting physicians to hypoxemia during anesthesia1,2, pneumonia3,4 and acute respiratory distress syndrome5. Currently, low and middle income countries are in need of an estimated one million pulse oximeters6, with only 41-70% of operating environments having access to a pulse oximeter in developing nations compared with 100% in developed nations7.To fill this need, previous groups have created lower cost pulse oximeters for use in low-income settings8–10.
In this article, we describe the design, calibration and validation of an open access pulse oximeter, known as the Glia pulse oximeter, at an estimated cost of under 25USD. The aim of the project is to provide free access to plans and software of the Glia pulse oximeter, allowing physicians around the world to build a Health Canada approved pulse oximeter that meets USA Food and Drug Administration (FDA) 501(k) certification for use in their clinic or operating room. To do so, the performance of the Glia pulse oximeter will be compared with arterial hemoglobin oxygen saturation (SaO2) as measured by blood gas analysis. In addition, performance testing comparing the calibrated Glia pulse oximeter to a standard Nellcor pulse oximeter will also be performed.

# Methods

Controlled Desaturation Study
The study was performed at Victoria Hospital, London Health Sciences Centre in London, Ontario. Before a desaturation procedure, volunteers were screened for exclusion criteria and informed written consent was obtained. FDA 501(k) protocols to test other pulse oximeter units was followed11.
Volunteers were laid flat with one arm raised onto a table and a radial arterial catheter was inserted for SaO2 measurements. The Glia oximeter and CO-oximeter was applied to adjacent fingers on contralateral arm. A Respiract was used to measure background end-tidal PCO2 and PO2 and a desaturation sequence was modified for those values.
At the beginning of the experiment, control oximeter data and two blood samples were drawn while the volunteer was breathing medical air. The volunteer then breathed 100% O2 and a maximum SpO2 was obtained. A hypoxic gas mixture containing nitrogen, oxygen and carbon dioxide was used to serially desaturate the volunteer using the Respiract to an SpO2 of 96-92%, 91-87%, 86-82%, 81-77%, 76-72% and 71-67%. Each plateau was allowed to stabilize for 30 seconds and N oximeter readings and arterial samples were taken at each plateau 20 seconds apart.

 1.	Walker, I. & Wilson, I. Anaesthesia in developing countries - a risk for patients. Lancet 968–969 (2009).
2.	Haynes, A., Weiser, T. & Berry, W. A surgical safety checklist to reduce morbidity and mortality in a global population. New Eng J Med 491–499 (2009).
3.	Duke, T., Graham, S. & Cherian, M. Oxygen is an essential medicine: a call for international action. Int J Tuberc Lung Dis 1362–1368 (2010).
4.	Ashraf, H., Jobayer, M. & Alam, N. Treatment of childhood pneumonia in developing countries.
5.	Kwizera, A. & Dunser, M. A Global Perspective on Acute Respiratory Distress Syndrome and the Truth about Hypoxia in Resource-limited Settings. Am. J. Respir. Crit. Care Med. 5–7 (2016).
6.	WHO. Global Pulse Oximetry Project. (2008).
7.	Funk, L., Weiser, T. & Berry, W. Global operating theatre distribution and pulse oximetry supply: an estimation from reported data. Lancet 1055–1061 (2010).
8.	Edwards, J. Taking the pulse of pulse oximetry in Africa. CMAJ 184, E244–E245 (2012).
9.	G, D. et al. Accuracy of the Lifebox pulse oximeter during hypoxia in healthy volunteers. Anesthesia 68, 1220–1223 (2013).
10.	L, F. et al. Evaluation of a large-scale donation of Lifebox pulse oximeters to non-physician anaesthetists in Uganda. Anaesthesia 69, 445–451 (2014).
11.	U.S. Food and Drug Administration. 501(k) clearances. http://www.fda.gov/MedicalDevices/ProductsandMedicalProcedures/DeviceApprovalsandClearances/510kClearances/default.html. Accessed July 27,2016.