01- Cover Letter & Executive Summary
Cover Letter

To Whom It May Concern:

Please find enclosed our application for a Medical Device Investigational Testing Authorization. Please do not hesitate to contact me at tarek@tarek.org or 519-488-6475 with questions regarding the application.

Regards,

Tarek Loubani, BSc (Hon), MD, CCFP(EM)
Associate Professor, Division of Emergency Medicine
Faculty of Medicine and Dentistry
Western University

Consultant, Department of Emergency Medicine
London Health Science Centre

Executive Summary

Pulse oximetry is an important tool used to assess a patient’s oxyhemoglobin saturation and it increases patient safety by alerting physicians to hypoxemia. Unfortunately, the costs of both the simplified and extended devices are prohibitive and reduce availability in smaller centres and poor countries, putting millions of patients in danger of easily treatable and preventable conditions. The technologies involved in the manufacture of a pulse oximeter are widely known and have no active patents covering them.

The Glia Inc. Pulse Oximeter is a device developed using current rapid prototyping technologies including 3D printing. The aim of Investigational Testing of the Glia Inc. Pulse Oximeter is to collect clinical data to be used in support of a medical device licence application to Health Canada with the goal of Health Canada approval of the manufacture and distribution of the Glia pulse oximeter. The aim of the project is to provide free and open access to plans and software of the Glia Inc. pulse oximeter, allowing physicians around the world to build a Health Canada approved pulse oximeter that meets USA Food and Drug Administration (FDA) 501(k) certification for use in their clinic or operating room.

This Investigational Testing Authorization application sets out the details of our proposed clinical performance testing of the Glia Inc. The Pulse Oximeter will be compared with arterial hemoglobin oxygen saturation (SaO2) as measured by blood gas analysis. Performance testing comparing the calibrated Glia pulse oximeter to the current gold standard pulse oximeter will also be performed.

For the investigation comparing the Pulse Oximeter with compared with arterial hemoglobin oxygen saturation as measured by blood gas analysis, the study design follows a standard protocol as recommended by the FDA. Study participants will have a radial arterial catheter inserted for SaO2 measurements. Both the Glia Inc. Pulse oximeter and the gold standard pulse oximeter will be applied to the patients fingers applied to adjacent fingers on contralateral arm. Under careful monitoring, oxygen saturation level will be decreased with gas mixtures of progressively decreasing oxygen and increasing nitrogen to a safe and pre-defined end point. Performance testing comparing the calibrated Glia pulse oximeter to the current gold standard pulse oximeter will involve contemporaneous measurements of oxygen saturation from both oximeters to enable direct comparison of the two oximeters.

Table of Contents
01…………………………………………………………Cover Letter & Executive Summary
02……………………………………………..……………………………….Table of Contents
03…………………………………………..………………………………….Application Form
04……………………………………………..………………Pre-submission Correspondence
05……………………………………………………………………………………Introduction
    05.01………………………..…………………………Manufacturer or Importer Identification
    05.02………………………………………………………………………Device Identification
    05.05……………………….…….……..Indications For Use, Intended Use, Contraindications
    05.06……………………………………………………………………...…...Device Labelling
07……………………………………………………..……………….Institutional Information
    07.01………………………………………………..………………..Investigator Information
    07.04…………………………………………….…………Name and Address of Institutions
    07.05……………………………………………………..…Research Ethics Board Approval
08…………………………………………………………………………….Study Documents
    08.01………………………………………………………………………………….Protocol
    08.03………………………………………………………………….Informed Consent Form
09……………………………………………………………………………………Appendices

03- Application Form
Application for Investigational Testing Authorization
(disponible en francais)
DEVICE CLASSIFICATION
	Class II	□	Class III	□	Class IV

DEVICE NAME (as it appears on label)
[Note: this is the device name for which the Authorization will be issued]
Pulse oximeter

PROTOCOL IDENTIFICATION:
Include the type of diagnosis or treatment for which the device will be sold.
The purpose of this device is to non-invasively measure oxyhemoglobin and deoxyhemoglobin, presenting clinicians and users with an oxygen saturation (SpO2), which is an estimate of partial pressure of oxygen in the bloodstream (pO2). The device is a monitoring device and has no specific therapeutic use.

NAME AND ADDRESS OF MANUFACTURER (as it appears on the label)
[Note: this is the name and address to which the Authorization will be issued]
Company Name	Glia Inc.
Street Address/P.O. Box: 54 Craig St. / City: London / Province/State: Ontario / Postal/Zip Code: N6C 1E8 / Country: Canada
Contact Name and Title:	Dr. Tarek Loubani, Primary investigator and Emergency Physician
Telephone No.: +1 519 488 6475		Fax No.: NA
E-Mail Address:	tarek@tarek.org

Application for Investigational Testing Authorization
(disponible en francais)
MAILING ADDRESS FOR REGULATORY CORRESPONDENCE (if different from 4)
Note:	(i) The Authorization will be issued to Company named in Item 4 but will be sent to the Company shown below if different.	(ii) The Company named below must be authorized by the manufacturer named in Item 4 to submit an authorization application on their behalf.
Company Name
Street Address/P.O. Box / City / Province/State / Postal/Zip Code / Country
Contact Name and Title:
Telephone No.:		Fax No.:
E-Mail Address:

DEVICE TYPE (check one only)
Single Device	✓
Medical Device Group
Medical Device Family
Medical Device Group Family
Test Kit
System

PREFERRED NAME CODE: (xxAAA) optional
IS THIS DEVICE A NEAR PATIENT IN VITRO DIAGNOSTIC (IVDD)?	Yes☐	No
IS THIS DEVICE INTENDED TO BE SOLD FOR HOME USE?	Yes☐	No
Application for Investigational Testing Authorization
(disponible en francais)
DEVICE USAGE CATEGORY
(73) Anaesthesiology			(84) Neurology
(74) Cardiovascular			(85) Obstetrics & Gynaecology
(76) Dental			(86) Ophthalmology
(77) Ear, Nose & Throat			(87) Orthopaedics
(78) Gastroenterology & Urology			(89) Physical Medicine
(79) General & Plastic Surgery			(90) Radiology/Imaging
(80) General Hospital	

FOR IVDDs ONLY
(75) Chemistry			(83) Microbiology
(81) Haematology			(88) Pathology
(82) Immunology			(91) Clinical Toxicology

DOES THIS DEVICE CONTAIN A DRUG?	Yes☐	No
(Note: this question does not apply to IVDDs)
# If yes
Brand /Trade Name of Drug:
Active Ingredient:
Drug Manufacturer:
Applicable Drug Identification Number (if any):

Application for Investigational Testing Authorization
(disponible en francais)
DEVICE DETAIL
Please provide the following information, where applicable for each component device, part or accessory.
Name of Device, Components, Parts and/or Accessories as per product label	Device Identification Number if previously assigned	Model or Catalogue number

Application for Investigational Testing Authorization
(disponible en francais)
ATTACHMENTS
In addition to items 1 to 11, of the Application for investigational testing, please indicate () which of the relevant information requirements listed below, are included as attachments to this application, or will be provided at a later date. For details regarding content and format, please refer to the guidance documents “Preparation of an Application for Investigational Testing - Medical Devices” and “Preparation of an Application for Investigational Testing - in Vitro Diagnostic Devices”
	Attached	To Come
Background Information	x
Risk Assessment
Ethics Committee or IRB Approval(s)	x
Protocol	x
Device Label	x
Investigator Agreements

If this Device contains a drug and it does not have a Drug Identification Number, I the Manufacturer of this device attest that the (☐ drug meets)	(☐ drug does not meet) acceptable standards of safety, efficacy and quality.
I hereby certify that the information provided on this application and in any attached documentation is correct, complete and in accordance with all relevant sections of the Medical Devices Regulations.
Name of Signing Official: Tarek Loubani
Signed:	Date: 2018 March 30

04- Pre-submission Correspondence

05- Introduction
05.01 Manufacturer or Importer Identification
Company Name	Glia Inc.
Street Address/P.O. Box: 54 Craig St. / City: London / Province/State: Ontario / Postal/Zip Code: N6C 1E8 / Country: Canada
Contact Name and Title:	Dr. Tarek Loubani, Primary investigator and Emergency Physician
Telephone No.: +1 519 488 6475		Fax No.: NA
E-Mail Address:	tarek@tarek.org

05.02 Device Identification
Device name: Pulse oximeter
Device label: See section 5.06- Device Labelling
Risk classification: This device is to be classified as a Class II device according to the classification rules set out in Schedule 1 of the Medical Devices Regulations
Units of device required for testing: 10 total (5 at Victoria Hospital, 5 at University Hospital). See 07.04-Name and Address of Institutions for more information.

05.05 Indications For Use, Intended Use, Contraindications
Indications for Use:
To measure arterial blood oxygen saturation.

Intended Use:
For use in adults over 18 years of age.

Contraindications:
None.

05.06 Device Labelling
The following device label (Figure 1) (designed according to requirements set out in Section 86 of the Medical Devices Regulations) will be included in the Operator's Manual, product monograph and will be affixed to the outer packaging of each investigational device. A similar label (Figure 2) containing identical elements (device name, manufacturer name, investigational statement) along with a unique investigational device code will be affixed directly to each device.

Glia Inc. Pulse Oximeter / INVESTIGATIONAL DEVICE / To be Used by Qualified Investigators Only INSTRUMENT DE RECHERCHÉ / Réservé uniquement à l'usage de chercheurs compétents

Figure 1. Device Label for Operator's Manual, product monograph and outer packaging

Figure 2. Sample device label to be affixed to each device

07- Institutional Information
07.01 Investigator Information
Tarek Loubani, MD. (PI) Department of Emergency Medicine, London Health Sciences
Centre, Western University
Alexander Pavlosky. Schulich School of Medicine and Dentistry, Western University

07.04 Name and Address of Institutions
Victoria Hospital Emergency Department
800 Commissioners Rd E
London, Ontario
N6A 5W9
519-685-8500

University Hospital Emergency Department
339 Windermere Road
London, Ontario
N6A 5A5
519 685-8500

07.05 Research Ethics Board Approval

08- Study Documents
08.01 Protocol
Protocol (Version: 2018.02, Date: 2018-01-30)

Background Information
Pulse oximetry is a vital tool in the rapid and non-invasive assessment of emergency patients, providing a continuous estimate of hemoglobin saturation in arterial blood. Some pulse oximeters are even capable of monitoring carboxyhemoglobin and methemoglobin, though these capabilities dramatically increase costs. Unfortunately, the costs of both the simplified and extended devices are prohibitive and reduce availability in smaller centres and poor countries, putting millions of patients in danger of easily treatable and preventable conditions. The technologies involved in the manufacture of a pulse oximeter are widely known and have no active patents covering them. This includes carboxy- and met-hemoglobin monitoring. Despite this, no high quality device exists whose design and construction are both research-validated and open to the public for study, production and dissemination. With current rapid prototyping technologies such as 3D printing, it is possible to create a very inexpensive pulse oximeter that meets or exceeds the gold standard. Design considerations include: cost, availability of parts, ease of construction, quality, ease of maintenance, ease of use. Preliminary engineering work has been carried out for the pulse oximeter, with design specifications set and a prototype built. The prototype was the result of collaboration between an engineer and an emergency physician with extensive local and international emergency experience. The pulse oximeter device emits light in four wavelengths: 590nm, 610nm, 660nm, 940nm. These allow for the calculation of oxygenated hemoglobin (660nm and 940nm), carboxyhemoglobin (610nm) and methemoglobin (590nm). The current bill of materials for the prototype is included in Appendix A. A printed circuit board of the prototype is included in Appendix B. A photograph of a completed prototype is included in Appendix C. Software has been designed and written to translate the pulse oximeter's sensor signals into usable values, display them, and provide appropriate alarms when necessary. As per widely-understood and validated concepts, an R value is calculated (ratio of emitted to absorbed light). This value is compared to a look-up table created during an empirical calibration phase, which assigns R values to oxygen saturation values. The value is then displayed for the user along with a graphical display of the pulsatile wave to ensure the quality of the signal. An appropriate user interface must accompany this software to allow for the safe setting of limits and alarms. The goal of this study is to validate and certify a pulse oximeter that measures hemoglobin, carboxyhemoglobin and methemoglobin. Once validated, this pulse oximeter will be certified with Health Canada, and then released under the Open Hardware License (OHL), such that hospitals and ministries of health in rural and impoverished communities in Canada and internationally would have easy access to these devices.

Hypothesis
The hypothesis of this study is that the experimental, 3D printed pulse oximeter will provide oxygen saturation readings with accuracy, equivalent to the gold standard pulse oximeter.

Objectives
The purpose of this 2-Phase study is to (1) calibrate a prototype pulse oximeter and (2) to conduct an equivalence test to determine if the prototype is equivalent to the gold standard pulse oximeter. The primary objectives are:
1) to compare arterial blood sample oxygen saturation measurements to the device's R value to create an empirical calibration curve as per the FDA document “Pulse Oximeters - Premarket Notification Submissions [510(k)s]: Guidance for Industry and Food and Drug Administration Staff” section 4.1.1 (http://www.fda.gov/RegulatoryInformation/Guidances/ucm341718.htm#s4);
2) to determine if the experimental pulse oximeter measurements correspond with the gold-standard pulse oximeter measurements;
and 3) to determine if the prototype pulse oximeter is equivalent to the gold standard pulse oximeter.

Method
This clinical trial consists of 2 Phases. Phase 1 will include 50 healthy, non-smoking adults (greater than 18 years of age) to calibrate the experimental, 3D printed pulse oximeter. Participants will be recruited through e-mail advertisements and posters between June 1 – August 2018. Subjects will participate in both the experimental arm and the control arm as both oximeters will be used to simultaneously generate data during the desaturation study (see Test Protocol). In Phase 2 we will conduct an equivalence study with 350 adult (greater than 18 years of age) emergency department (ED) patients visiting either Victoria or University Hospital at London Health Sciences Centre (LHSC). Investigators will approach patients at the invitation of a member in the circle of care to discuss the study. Recruitment will occur between June 1 and August 30, 2018. After going through the LOI and gaining documented written consent, the patient will be enrolled. All study procedures will be conducted in the ED (Victoria and University Hospital) at LHSC by the PI.

All study procedures will be conducted in the emergency department (VH and UH) at LHSC with an emergency-trained physician (EP) present at all times (PI). The estimated time frame for the duration of the study, including the enrolment phase and treatment phase, is from June 1, 2018 and December 31, 2019. This study does not require any patient follow-up.

Phase 1 (Calibration; approximately 50 participants; Sept 1, 2018 – April 30, 2019):
Upon obtaining informed consent, participants will be positioned supine in a hospital bed. Participants will be cannulated with an indwelling arterial catheter for frequent sampling of arterial blood. The experimental pulse oximeter will be placed on the participant's finger and continuously monitored. A gold-standard pulse oximeter (GE Carescape B850 Patient Monitor; license 119340) will be placed on a separate finger on the participant's same hand and continuously monitored as per the institution's procedural sedation protocols. Participant samples will be tested with a CO-oximeter to ensure that background carboxyhemoglobin (COHb) is less than 2%. Two blood samples will be taken with the participant on room air. The participant will then be fitted with 100% oxygen by non-rebreather mask. Under careful monitoring by the Emergency Physician, oxygen saturation level will be decreased with gas mixtures of progressively decreasing oxygen and increasing nitrogen. Measurements will be recorded until the participant has unstable vitals (e.g., hypotension, tachycardia), is symptomatic or a pulse saturation of 70% is achieved. Oxygen levels will then be immediately restored. Arterial blood samples will be drawn and oxygen saturation will be measured when the gas mixture and R value readings have stabilized. 2 cubic centimetres of blood will be drawn for every 5% drop in blood oxygen saturation from 100% to 70% or earlier if the participant experiences unstable vitals or is symptomatic (Table 1). Phase 1 of the study will take approximately 60 minutes to complete.
Step	Target SpO2	PetCO2	Target PetO2	G1 flow
1	95-99	40	90	35
2	95-99	40	59	35
3	90-94	40	50	35
4	85-89	40	44	35
5	80-84	40	40	35
6	75-79	40	37	35
7	95-99	40	100	35
8	100	40	500	35

Table 1: Respiract desaturation protocol

This study design is routine for all pulse oximeters, and is as recommended by the FDA guidance document “Pulse Oximeters - Premarket Notification Submissions [510(k)s]: Guidance for Industry and Food and Drug Administration Staff”, section 4.1.1 (http://www.fda.gov/RegulatoryInformation/Guidances/ucm341718.htm#s4 )

Phase 2 (validation; approximately 350 ED patients):
Patients who have attended the ED for care are eligible to participate. Investigators will approach patients at the invitation of a member in the circle of care to discuss the study. Recruitment will occur between May 1 and December 31, 2019. After going through the LOI and gaining documented written consent, the volunteer/patient will be enrolled. The experimental and control pulse oximeters will be placed on two different fingers on the same hand. Participants will not be cannulated in Phase 2. Patients will then have one measurement taken with the experimental pulse oximeter and another measurement taken with a gold standard pulse oximeter. Measurements of oxygen saturation will be recorded to establish equivalence of the two oximeters. Part two of the study takes 5 minutes to complete, is completely non-invasive, and will not interfere with regular clinical care. There will be no interference with a patient's normal care during this phase.

Medical Device: Experimental Oximeter
The experimental oximeter will be manufactured by the research team in London, Ontario under the supervision of Glia Inc. (Corporation #924742-4; HC Company ID #141507). A bill of materials can be found in Table 2. The oximeter has not been granted authorization for use in Canada.
Mfr. #	Manufacturer	Description	Order Qty.
VJ0603V105MXQCW1BC	Vishay	Multilayer Ceramic Capacitors MLCC - SMD/SMT 0603 1uF 10volts Y5V 20%	9
GRM188R60J475KE19D	Murata	Multilayer Ceramic Capacitors MLCC - SMD/SMT 0603 4.7uF 6.3volts X5R 10%	2
06036D106MAT2A	AVX	Multilayer Ceramic Capacitors MLCC - SMD/SMT 6.3V 10uF X5R 0603 20% Tol HIGH CV	2
VJ0603Y104JXJPW1BC	Vishay	Multilayer Ceramic Capacitors MLCC - SMD/SMT 0603 0.1uF 16volts X7R 5%	4
MCP1700T-3002E/TT	Microchip	LDO Voltage Regulators 250mA Adj LDO 2%	1
CRCW06032K00FKEA	Vishay	Thick Film Resistors - SMD 1/10watt 2.0Kohms 1%	1
CRCW0603590KFKEA	Vishay	Thick Film Resistors - SMD 1/10watt 590Kohms 1%	1
CRCW02011M00FKED	Vishay	Thick Film Resistors - SMD 1/20watt 1Mohms 1% 100ppm	3
FSMSM	TE Connectivity	Tactile Switches 3.5X6 SMT TACT TACT SWITCH	2
MCP1640T-I/CHY	Microchip	Voltage Regulators - Switching Regulators 500 kHz 300 mA Syn. PWM/PFM enabled	1
AT-2440-TWT-R	PUI Audio	Speakers & Transducers 3V 80dBA 4000Hz	1
744028004	Wurth Electronics	Fixed Inductors WE-TPC 2811 4.7uH .85A .265Ohm	1
STM32F030F4P6	STMicroelectronics	ARM Microcontrollers - MCU Value-Line ARM MCU 16kB 48 MHz	1
3-1734592-0	TE Connectivity	FFC & FPC Connectors 0.5mm PITCH B/C 30P	1
XG8S-0231	Omron	Headers & Wire Housings Connector	25
SML-LX0603SRW-TR	Lumex	Standard LEDs - SMD Super Red, 660nm 1.7V, 45mcd	1
SML-LX0603SYW-TR	Lumex	Standard LEDs - SMD Super Yellow, 590nm 2V, 60mcd	1
598-8030-107F	Dialight	Standard LEDs - SMD Orange Water Clr 150mcd 610nm	1
APT1608F3C	Kingbright	Infrared Emitters IR 940nm 120 deg Water Clr 1.2 mW/sr	1
TSL13T	ams	Light To Frequency & Light To Voltage Light to Voltage Converter	1

Table 2: Bill of Materials of experimental oximeter.

Subjects
According to FDA recommendations, 15 participants is the minimum number required for calibration studies. In order to ensure the experimental pulse oximeter is adequately calibrated, we have determine a sample of 50 participants is sufficient for Phase 1. In Phase 2, to maintain a power of 0.80 with an alpha of 0.05, the sample size was calculated at 350 participants to establish equivalence. See sample size calculations below.
### Phase 1 sample size
The phase 1 (calibration) sample size was determined by considering the FDA and ISO recommendations for calibration studies on healthy volunteers. While 15 is the minimum number required (See FDA document attached), we expect a sample size of 30 patients. To allow for unexpected variability, we have requested a sample size of 50 from the REB, though this number might be reduced after a interim analysis at 25-30 patients.
### Phase 2 Sample Size
The phase 2 sample size was calculated in the following way: Consider a clinical trial comparing two groups, a standard therapy (s) and a novel therapy (n). In each group, a proportion of subjects responds to the treatment: Ps and Pn. If the intention of the study is to show that the two groups are equivalent, the usual formulation of the null hypothesis (Ps=Pn) encounters logical difficulty. A statistical test may fail to reject this null hypothesis, but this will not mean that the two treatments are equivalent.
Blackwelder (Controlled Clinical Trials 1982; 3: 345-353) proposes a solution. If a difference between the two treatments, call it D, is specified that practically represents equivalence, then the null hypothesis can be restated to include the specified difference. In other words, that: Ps is greater than or equal to Pn + D. Rejection of this hypothesis implies that the difference between the standard and novel treatments is less than or equal to D, indicating equivalence.
The sample size needed to reject this hypothesis at alpha = 0.05 and beta = 0.20 (power of 80%) is:
(Z 0.95 + Z 0.80)2 [Ps(1-Ps) + Pn(1-Pn)] / (Ps-Pn-D)2
We made the following assumptions:
1) "response rate" being the proportion of correct readings, assuming a 4% variance is .96 in both the "gold standard" and "new pulse ox" groups.
2) A practical difference (accepted clinical difference between the novel and standard device) of also 4%.
With these numbers, we require 296 patients/volunteers in both groups. Equivalence is hard to prove and has the highest sample size needed. We rounded this number up to 350 to account for unexpected variability or difficulties.
If we desired a non-inferiority study, those numbers go down significantly to 13/group. If there is truly no difference between the standard and experimental treatment, then 26 patients are required to be 80% sure that the lower limit of a one-sided 95% confidence interval (or equivalently a 90% two-sided confidence interval) will be above the non-inferiority limit of -4. However, because of the very short amount of time required for each participant and no risk, we elected to run an equivalence trial.

Inclusion Criteria
This study has 2 Phases. In Phase 1, healthy, non-smoking adult volunteers greater than 18 years of age are invited to participate. In Phase 2, all adult (>18) non-critical patients visiting the ED at VH and UH are eligible to participate.

Exclusion Criteria
Potential participants will be excluded based on the following criteria: Phase 1: greater than 70 years of age, current smokers, previous history of cardiac disease (e.g., atrial fibrillation, myocardial infarction, congestive heart failure), history of moderate or severe asthma, history of seizures, history of stroke or transient ischemic attack, and pregnant women will be excluded. Patients with any active pulmonary disease such as pneumonia will be excluded. Individuals who are unwilling to participate or are less than 18 years old will be excluded, as will those who are unable to consent. Phase 2, critical patients will be excluded from participating. Individuals who are unwilling to participate or are less than 18 years old will be excluded, as will those who are unable to consent.

Stopping rules
Subject wishes to withdraw from the protocol.

Data Analysis
Data will be entered directly into a study-specific LibreOffice database by investigators (The Document Foundation, Berlin, Germany). No personal identifiers will be collected in this study. Standard descriptive statistics will be summarized using means and standard deviations and differences in proportions will be assessed by the Pearson chi-squared statistic where appropriate. All data analyses will be performed using R (R Foundation, Vienna, Austria) by investigators.

Phase 1: to validate the experimental pulse oximeter, readings from the arterial blood sample will be plotted against the R value from the prototype to establish a best-fit curve.

Phase 2: two one sided tests will be employed to determine if the experimental pulse oximeter is equivalent to the gold standard. The confidence interval around the mean difference between the prototype and gold standard will be compared to predetermined acceptance criterion.

Risks and harms
During the study protocol in Phase 1, participants will be cannulated with an indwelling arterial catheter for frequent sampling of arterial blood. Participants may feel discomfort during the catheter insertion process. The participant will also be fitted with a non-rebreather mask as part of the Phase 1 study protocol in order to reduce the oxygen saturation level with gas mixtures of progressively decreasing oxygen and increasing nitrogen. Participants may experience low blood pressure, tachycardia or chest pain. Phase 1 is estimated to take 60 minutes to complete. In Phase 2 of the study, participants will wear both the experimental and gold standard pulse oximeter. There are no additional risks associated with wearing two pulse oximeters for the five minute duration of the study.
The entirety of the study protocol will be conducted within the emergency department at LHSC, on a hospital bed, with an emergency trained physician monitoring at all times. In Phase 1 the arterial catheter will be inserted by an experienced emergency physician. Should an adverse event occur, the participant will be monitored one-on-one by the physician and treated according to best medical practice. Arterial oxygen saturation will be continuously monitored as per the institution's procedural sedation protocols.

Informed consent
Participants will be provided with the LOI and informed consent document. The investigator will verbally describe the study to the participant. The participant will be made aware of the purpose of the study, what their participation in the study entails, assured there is no deception in the protocol, be made aware of the risks of participating, and be assured that they may end their participation in the study at any time without consequence. The investigator will provide ample time to read the LOI and actively solicit any questions or concerns. Documented written consent will be obtained before enrolling the participant. For Phase 2, patients will have first contact within the circle of care for participation.

08.03 Informed Consent Form
Version Date: April 06, 2016

Division of Emergency Medicine

Title of Project: Calibration and Validation of Low-Cost 3D Printed Pulse Oximeter

Principal Investigator: Dr. Tarek Loubani, MD, CCFP (EM)

Co-Investigators: Dr. Melanie Columbus, PhD, Mrs. Kristine Van Aarsen, MSc

Background & Purpose

You are invited to participate in a research study to calibrate a low-cost 3D printed pulse oximeter. A pulse oximeter is a sensor device placed on the finger to measure oxygen levels in the blood. All emergency department patients and adults over the age of 18 are eligible to participate in this study. This letter contains information to help you decide whether or not to participate in this research study. It is important for you to understand why the study is being conducted and what it will involve. Please take the time to review it carefully and feel free to ask questions if anything is unclear or there are words/phrases that you do not understand.

The purpose of this study is to calibrate and validate a low-cost 3D printed pulse oximeter that measures oxygen levels in the blood. Measurements from the experimental pulse oximeter will be compared to a gold standard – the pulse oximeter currently used in the emergency department at London Health Sciences Centre (LHSC). The experimental pulse oximeter can be manufactured at a fraction of the cost of currently employed devices and may be a cost-effective alternative for hospitals and clinics in both the developed and developing worlds.

You are being asked to help validate a new low-cost 3D printed pulse oximeter. Your participation in the study involves completion of the study protocol. All study procedures will be conducted in the emergency department at Victoria Hospital and University Hospital at LHSC by an experienced emergency physician (EP). 350 participants will be recruited examine the validity of the calibrated experimental pulse oximeter. The experimental and control pulse oximeters will be placed on two different fingers on the same hand. Measurements of oxygen saturation will be recorded to establish equivalence of the two oximeters. This study takes approximately 5 minutes to complete, is completely non-invasive, and will not interfere with regular clinical care.

This study is part of an unfunded research project being under taken by Dr. Tarek Loubani of Western University. Should you require any further information, please feel free to contact: Dr. Tarek Loubani (tlouban@uwo.ca), Carrie Wakem (carrie.wakem@lhsc.on.ca) or Melanie Columbus (melanie.columbus@lhsc.on.ca) at 519-685-8500 ext 55014.

Risks & Benefits

There are no additional risks associated with wearing two pulse oximeters for the five minute duration of the study. You may not benefit personally from the study. Participation in this study will not impact your clinical course, academic status or employment. Results of this study may benefit society by validating a low-priced pulse oximeter whose design is freely available which may impact emergency rooms in both the developed and developing worlds. You will not be compensated for your participation in this study.

Participation in this study is voluntary. You may refuse to participate or withdraw at any point in the study without impact on your clinical course, employment or academic status. If you withdraw, any data collected may also be withdrawn at your request.

Confidentiality

Your results will be stored in a locked cabinet in a secure office, will be viewed only by members of the research team, and will be destroyed at the completion of this study. If the results of this study are published, your name will not be used and no information that discloses your identity will be released or published. You do not waive any legal rights by participating in this study. Data will be retained for a period of 15 years after publication in a secure place, after which time it will be disposed of in a secure manner (e.g. shredded or electronically deleted).

Qualified representatives of the following organizations may look at the study records at the site where these records are held, for quality assurance (to check that the information collected for the study is correct and follows proper laws and guidelines).
Examples include:
- Representatives of Lawson Quality Assurance Education Program
- Representatives of the University of Western Ontario Health Sciences Research Ethics Board that oversees the ethical conduct of this study

If you have any questions about your rights as a research participant or the conduct of the study you may contact Dr. David Hill, Scientific Director, Lawson Health Research Institute, (519) 667-6649

Should you require any further information, please feel free to contact: Dr. Tarek Loubani (tlouban@uwo.ca), Carrie Wakem (carrie.wakem@lhsc.on.ca) or Melanie Columbus (melanie.columbus@lhsc.on.ca) at 519-685-8500 ext 55014.

Thank you,

Dr. Tarek Loubani, MD, CCFP (EM)
Assistant Professor
Schulich School of Medicine and Dentistry
Western University

Division of Emergency Medicine

Consent to Participate

Title of Project: Calibration and Validation of Low-Cost 3D Printed Pulse Oximeter

Principal Investigator: Dr. Tarek Loubani, MD, CCFP (EM)

Co-investigator: Dr. Melanie Columbus, PhD, Mrs. Kristine Van Aarsen, MSc

I have read the Letter of Information, have had the nature of the study explained to me and I agree to participate. All questions have been answered to my satisfaction.

___________________	________________________	____________
Participant Signature	Participant Name (Printed)	Date

___________________	________________________	___________
Signature of Person	Person Obtaining Informed	Date
Obtaining Informed Consent	Consent (Printed)

09- Appendices